## Slide 1: Measurement of Emission Color of Single Molecules
Taehwan Kim
3/23/17

## Slide 2: Single-Molecule Color Sensing Methods

## Slide 3: Single-Molecule Color Sensing Methods
GOAL: Sensing spectral mean of a single emitter along with its position
Available Options
  Optical filter bank & λ stack construction
  PSF engineering based on chromatic dispersion (gratings1 , prism2 , programmable phase mask with 4f setup3 , etc.)
What is desired?
1) Ma, Y. et al., High-Throughput Single-Molecule Spectroscopy in Free Solution. Anal. Chem. (2000).
2) Zhang, Z. et al., Ultrahigh-throughput single-molecule spectroscopy and spectrally resolved super-resolution microscopy. Nat. Meth. (2015).
3) Shechtman, Y. et al., Multicolour localization microscopy by point-spread-function engineering. Nat. Photon. (2016).
Notes: S_i(f) = \frac{{i_0}^2}{2}\left[ \frac{{\tau}_c/(2\pi)}{1+\left(\pi f {\tau}_c \right)^2} \left\{ 1- e^{-\frac{2\tau}{{\tau}_c}} \left[ \cos(2\pi f\tau) + \frac{1}{\pi f \tau_c} \sin(2\pi f \tau) \right] \right\} \right]
\tau_c = \frac{2}{\textrm{linewidth}},\quad i_0: \textrm{nominal photocurrent amplitude},\quad \tau: \textrm{path delay}

## Slide 4: Color Sensing Methods: Criteria
Minimum overhead in the optical setup
  Less aberration & insertion loss, easy to apply, avoid complicated calibration
Minimum impact on lateral/axial localization accuracy
  Localization accuracy degrades by √2 per imaging path split
Avoid PSF footprint enlargement
  In SMLM context, minimize the probability of overlap of PSFs (impacts temporal resolution, especially in 2D samples like focal adhesion)
Flexibility (no need for a priori wavelength info)
Good λ estimation accuracy for given photon budget
Avoid possibility of bias from non-idealities
Notes: S_i(f) = \frac{{i_0}^2}{2}\left[ \frac{{\tau}_c/(2\pi)}{1+\left(\pi f {\tau}_c \right)^2} \left\{ 1- e^{-\frac{2\tau}{{\tau}_c}} \left[ \cos(2\pi f\tau) + \frac{1}{\pi f \tau_c} \sin(2\pi f \tau) \right] \right\} \right]
\tau_c = \frac{2}{\textrm{linewidth}},\quad i_0: \textrm{nominal photocurrent amplitude},\quad \tau: \textrm{path delay}

## Slide 5: SM color Sensing From PSF Fitting
Diffraction-limited image of a single molecule already contains color information
In principle, it should be possible to estimate emitter color from the scale of the fitted PSF
Spectrally resolved SMLM: just installation of a ImageJ plugin!
  Literally zero impact on the optical setup/xyz accuracy
  No PSF modification (more than what’s required for 3D)
  No a priori information
Notes: S_i(f) = \frac{{i_0}^2}{2}\left[ \frac{{\tau}_c/(2\pi)}{1+\left(\pi f {\tau}_c \right)^2} \left\{ 1- e^{-\frac{2\tau}{{\tau}_c}} \left[ \cos(2\pi f\tau) + \frac{1}{\pi f \tau_c} \sin(2\pi f \tau) \right] \right\} \right]
\tau_c = \frac{2}{\textrm{linewidth}},\quad i_0: \textrm{nominal photocurrent amplitude},\quad \tau: \textrm{path delay}

## Slide 6: Single-Molecule Color Sensing Methods
Questions to be answered
  Good enough λ estimation accuracy in realistic scenario?
    Axial position dependency in 2D situation
  Possibility of bias from non-idealities (not limited to)
    Chromatic/spherical aberration
    Spectral profile of the fluorophore (bandwidth, skewness)
    Non-uniform background
    Deconvolution artifacts
    Pixelation
    Motion blur
Notes: S_i(f) = \frac{{i_0}^2}{2}\left[ \frac{{\tau}_c/(2\pi)}{1+\left(\pi f {\tau}_c \right)^2} \left\{ 1- e^{-\frac{2\tau}{{\tau}_c}} \left[ \cos(2\pi f\tau) + \frac{1}{\pi f \tau_c} \sin(2\pi f \tau) \right] \right\} \right]
\tau_c = \frac{2}{\textrm{linewidth}},\quad i_0: \textrm{nominal photocurrent amplitude},\quad \tau: \textrm{path delay}

## Slide 7: Theoretical Performance Analysis: Accuracy

## Slide 8: Performance of Measurement Method
Task
  Estimation of x, y, z, avg(λ) of a single emitter from values of sensor array (EMCCD, sCMOS) within certain ROI
A “measurement method” includes:
  Optical setting (affecting SNR, bias)
  PSF design (encoding scheme: prism/grating, astigmatism, double helix, etc.)
  Reconstruction algorithm (LS/MLE, ROI selection, b/g correction, etc.)
Performance metrics
  Highest accuracy (smallest σ) for given number of photons & noise profile
  Unbiased
  PSF footprint
Notes: S_i(f) = \frac{{i_0}^2}{2}\left[ \frac{{\tau}_c/(2\pi)}{1+\left(\pi f {\tau}_c \right)^2} \left\{ 1- e^{-\frac{2\tau}{{\tau}_c}} \left[ \cos(2\pi f\tau) + \frac{1}{\pi f \tau_c} \sin(2\pi f \tau) \right] \right\} \right]
\tau_c = \frac{2}{\textrm{linewidth}},\quad i_0: \textrm{nominal photocurrent amplitude},\quad \tau: \textrm{path delay}

## Slide 9: Design of Measurement Method
Major design decisions in our context
  PSF design/PSF model to use (optical setting is determined by this)
  Deconvolution algorithm
Today: theoretical bound of λ accuracy/impact of PSF choice
  Fix other factors
    Magnification, noise, pixel size, ROI, NA, illumination, emitter spectra BW
  Consider only major physical implication of PSF design choice (e.g. consider the loss of photons when beam-split is necessary, but not the insertion loss from additional optical elements)
  Ignore aberration (for now)

## Slide 10: Problem Setting
PSF options: sensitivity is PSF model dependent
  Plain-vanilla circular aperture
  Multifocus (sum of the Fisher matrix from each image plane1)
  Astigmatism
  Double-helix
Assumptions
  NA=1.4, effective pixel size=100nm, ROI=20x20 pixels
  Only shot noise, uniform background of 10 photons/pixel
Procedure
  Compute CRLB for different z location & number of photons
1) Ram, S. et al., High Accuracy 3D Quantum Dot Tracking with Multifocal Plane Microscopy for the Study of Fast Intracellular Dynamics in Live Cells. Biophysical Journal (2008).
Notes: S_i(f) = \frac{{i_0}^2}{2}\left[ \frac{{\tau}_c/(2\pi)}{1+\left(\pi f {\tau}_c \right)^2} \left\{ 1- e^{-\frac{2\tau}{{\tau}_c}} \left[ \cos(2\pi f\tau) + \frac{1}{\pi f \tau_c} \sin(2\pi f \tau) \right] \right\} \right]
\tau_c = \frac{2}{\textrm{linewidth}},\quad i_0: \textrm{nominal photocurrent amplitude},\quad \tau: \textrm{path delay}

## Slide 11: Theoretical Performance Evaluation
For given setting, accuracy can be estimated using Fisher information/Cramér–Rao lower bound (CRLB)
  Fisher information matrix (FI): sensitivity of an observable quantity (e.g. PSF) to changes in its underlying parameters (e.g. emitter position)
  Cramér-Rao lower bound: theoretical best-case precision with any unbiased estimator given a PSF model and a noise model
    CRLB = diagonal elements of the inverse of FI matrix
  Depending on the algorithm, actually achievable; not just fictitious number
    Empirically, MLE yields close result to CRLB in most cases (although Insight3 uses LS)
  Common metric to compare the performance of different PSFs1 or optimize PSF for best precision2
1) Badieirostami, M. et al., Three-dimensional localization precision of the double-helix point spread function versus astigmatism and biplane. Appl. Phys. Lett (2010).
2) Shechtman, Y. et al., Optimal Point Spread Function Design for 3D Imaging. Phys. Rev. Lett. (2014).
Notes: \begin{gather*}
I_{ij}(\boldsymbol{\theta})=-E\left(\frac{\partial^2 \ln (f(s;\boldsymbol{\theta}))}{\partial\theta_i\partial\theta_j}\right)\\\boldsymbol{\theta}:\textsf{parameters being estimated},\: f(s;\boldsymbol{\theta}):\textsf{probability of $s$ being the measurement given $\boldsymbol{\theta}$}
\end{gather*}

## Slide 12: Cramér–Rao Lower Bound
Fisher Information Matrix
Cramér–Rao lower bound
Notes: \begin{gather*}
I_{ij}(\boldsymbol{\theta})=-E\left(\frac{\partial^2 \ln \left( f(s;\boldsymbol{\theta}) \right)}{\partial\theta_i \partial\theta_j}\right)\\
\boldsymbol{\theta}\textsf{: the vector of parameters being estimated $(x_0, y_0, z_0, \lambda, N_{photon}, \beta)$, }\\
\beta \textsf{: the number of background photons per pixel}\\
f(s;\boldsymbol{\theta})\textsf{: probability of $s$ being the measurement (i.e. PSF image) given $\boldsymbol{\theta}$}
\end{gather*}
\left[(\mathbf{I}(\boldsymbol{\theta}))^{-1}\right]_{ii}=\textrm{CRLB}_i

## Slide 13: Cramér–Rao Lower Bound
Explicit form in the presence of pixelation & shot noise only1
1) Ober, R. J et al., Localization Accuracy in Single-Molecule Microscopy. Biophysical Journal (2004).
Notes: \begin{gather*}
I_{ij}(\boldsymbol{\theta})=\sum_{k=1}^{N_{pixel}}{\frac{1}{\mu_k(\boldsymbol{\theta})+\beta}\left(\frac{\partial \mu_k(\boldsymbol{\theta})}{\partial \theta_i}\right)\left(\frac{\partial \mu_k(\boldsymbol{\theta})}{\partial \theta_j}\right)}\\
\mu_k(\boldsymbol{\theta})\textsf{: the value of the model PSF in pixel k}\\
N_{pixel}\textsf{: the number of pixels in the ROI}
\end{gather*}
\begin{gather*}
\mu_k(\boldsymbol{\theta})=\int_{\mathscr{D}_k}{dxdy \left( Nq(\vec{r}-\vec{r}_0) \right)}\\
q(\vec{r}\,;z_0,\lambda)\textsf{: normalized PSF model,} \: a\textsf{: pixel size}\\
\mathscr{D}_k\textsf{: $k$th pixel area,}\: \vec{r}=(x,y)\textsf{: image plane location,}\:\vec{r}_0=(x_0,y_0)\textsf{: lateral emitter location}
\end{gather*}
\begin{gather*}
\frac{\partial\mu_k}{\partial N}=\int_{\mathscr{D}_k}{dxdy q(\vec{r}-\vec{r}_0)}\\
\frac{\partial\mu_k}{\partial\theta}=2N\int_{\mathscr{D}_k}{dxdy\,\frac{\partial q(\vec{r}-\vec{r}_0)}{\partial \theta}},\:\textsf{For $\theta=x_0, y_0, z_0, \lambda$}
\end{gather*}

## Slide 14: CRLB: Circular Aperture Case
3D PSF model for simple circular aperture: Born & Wolf model
Notes: \begin{gather*}
q_{z_0}(\vec{r}-\vec{r}_0\,)=\left\|A \int_{0}^{1}{J_0\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}-\vec{r}_0\| \rho\right)e^{-j\frac{\pi}{\lambda}\rho^2 z_0\left(\frac{\textrm{NA}^2}{n_i}\right) }\rho d\rho} \right\|^2\\
A\textsf{: normalization constant,}\: n_i\textsf{: refractive index of the immersion oil}\\
J_0\textsf{: zeroth order Bessel function of the first kind}
\end{gather*}
\begin{gather*}
q_{z_0}(\vec{r}-\vec{r}_0\,)=\frac{1}{A_{z_0}}\left(u_{z_0}^2(\vec{r}-\vec{r}_0\,) + v_{z_0}^2(\vec{r}-\vec{r}_0\,)\right)\\
u_{z_0}(\vec{r}-\vec{r}_0\,)=\mathlarger{\int}_{0}^{1}{J_0\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}-\vec{r}_0\| \rho\right)\cos\left({\frac{\pi}{\lambda}\rho^2 z_0 \left(\frac{\textrm{NA}^2}{n_i}\right)}\right)\rho d\rho} \\
v_{z_0}(\vec{r}-\vec{r}_0\,)=\mathlarger{\int}_{0}^{1}{J_0\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}-\vec{r}_0\| \rho\right)\sin\left({\frac{\pi}{\lambda}\rho^2 z_0 \left(\frac{\textrm{NA}^2}{n_i}\right)}\right)\rho d\rho} \\
A_{z_0} = \int_{\mathbb{R}^2} {\left(u^2(\vec{r}\,) + v^2(\vec{r}\,)\right)}d\vec{r}
\end{gather*}

## Slide 15: CRLB: Circular Aperture Case
Notes: \begin{gather*}
\frac{\partial q(\vec{r}-\vec{r}_0\,)}{\partial x_0} = \frac{2}{A_{z_0}} \left( u(\vec{r}-\vec{r}_0\,) \frac{\partial u(\vec{r}-\vec{r}_0\,)}{\partial x_0} + v(\vec{r}-\vec{r}_0\,)\frac{\partial v(\vec{r}-\vec{r}_0\,)}{\partial x_0} \right)\\
\frac{\partial q(\vec{r}-\vec{r}_0\,)}{\partial y_0} = \frac{2}{A_{z_0}} \left( u(\vec{r}-\vec{r}_0\,) \frac{\partial u(\vec{r}-\vec{r}_0\,)}{\partial y_0} + v(\vec{r}-\vec{r}_0\,)\frac{\partial v(\vec{r}-\vec{r}_0\,)}{\partial y_0} \right)\\
\frac{\partial q(\vec{r}-\vec{r}_0\,)}{\partial \lambda} = \frac{2}{A_{z_0}} \left( u(\vec{r}-\vec{r}_0\,) \frac{\partial u(\vec{r}-\vec{r}_0\,)}{\partial \lambda} + v(\vec{r}-\vec{r}_0\,)\frac{\partial v(\vec{r}-\vec{r}_0\,)}{\partial \lambda} \right) - \frac{1}{A_{z_0}}\frac{\partial A_{z_0}}{\partial \lambda}q(\vec{r}-\vec{r}_0\,)
\end{gather*}
\begin{gather*}
\frac{\partial u(\vec{r}\,)}{\partial x_0}=\frac{2\pi}{\lambda}\textrm{NA}\frac{x}{\|\vec{r}\,\|} \int_{0}^{1}{J_1\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}\,\|\rho \right)\cos\left({\frac{\pi}{\lambda}\rho^2 z_0 \left(\frac{\textrm{NA}^2}{n_i}\right)}\right)\rho^2 d\rho} \\
\frac{\partial v(\vec{r}\,)}{\partial x_0}=\frac{2\pi}{\lambda}\textrm{NA}\frac{x}{\|\vec{r}\,\|} \int_{0}^{1}{J_1\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}\,\|\rho \right)\sin\left({\frac{\pi}{\lambda}\rho^2 z_0 \left(\frac{\textrm{NA}^2}{n_i}\right)}\right)\rho^2 d\rho} \\
\frac{\partial u(\vec{r}\,)}{\partial y_0}=\frac{2\pi}{\lambda}\textrm{NA}\frac{y}{\|\vec{r}\,\|} \int_{0}^{1}{J_1\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}\,\|\rho \right)\cos\left({\frac{\pi}{\lambda}\rho^2 z_0 \left(\frac{\textrm{NA}^2}{n_i}\right)}\right)\rho^2 d\rho} \\
\frac{\partial v(\vec{r}\,)}{\partial y_0}=\frac{2\pi}{\lambda}\textrm{NA}\frac{y}{\|\vec{r}\,\|} \int_{0}^{1}{J_1\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}\,\|\rho \right)\sin\left({\frac{\pi}{\lambda}\rho^2 z_0 \left(\frac{\textrm{NA}^2}{n_i}\right)}\right)\rho^2 d\rho} \\
\end{gather*}

## Slide 16: CRLB: Circular Aperture Case
Notes: \begin{gather*}
\frac{\partial q(\vec{r}-\vec{r}_0\,)}{\partial x_0} = \frac{2}{A} \left( u(\vec{r}-\vec{r}_0\,) \frac{\partial u(\vec{r}-\vec{r}_0\,)}{\partial x_0} + v(\vec{r}-\vec{r}_0\,)\frac{\partial v(\vec{r}-\vec{r}_0\,)}{\partial x_0} \right)\\
\frac{\partial q(\vec{r}-\vec{r}_0\,)}{\partial y_0} = \frac{2}{A} \left( u(\vec{r}-\vec{r}_0\,) \frac{\partial u(\vec{r}-\vec{r}_0\,)}{\partial y_0} + v(\vec{r}-\vec{r}_0\,)\frac{\partial v(\vec{r}-\vec{r}_0\,)}{\partial y_0} \right)\\
\frac{\partial q(\vec{r}-\vec{r}_0\,)}{\partial \lambda} = \frac{2}{A} \left( u(\vec{r}-\vec{r}_0\,) \frac{\partial u(\vec{r}-\vec{r}_0\,)}{\partial \lambda} + v(\vec{r}-\vec{r}_0\,)\frac{\partial v(\vec{r}-\vec{r}_0\,)}{\partial \lambda} \right) - \frac{1}{A}\frac{\partial A}{\partial \lambda}q(\vec{r}-\vec{r}_0\,)
\end{gather*}
\begin{gather*}
\frac{\partial u(\vec{r}\,)}{\partial x_0}=\frac{2\pi \textrm{NA}}{\lambda}\frac{x}{\|\vec{r}\,\|} \int_{0}^{1}{J_1\left(\frac{2\pi \textrm{NA}}{\lambda} \rho \|\vec{r}\,\|\right)\cos\left({\frac{\pi \textrm{NA}^2 z_0}{\lambda n_m}\rho^2}\right)\rho^2 d\rho} \\
\frac{\partial v(\vec{r}\,)}{\partial x_0}=\frac{2\pi \textrm{NA}}{\lambda}\frac{x}{\|\vec{r}\,\|} \int_{0}^{1}{J_1\left(\frac{2\pi \textrm{NA}}{\lambda} \rho \|\vec{r}\,\|\right)\sin\left({\frac{\pi \textrm{NA}^2 z_0}{\lambda n_m}\rho^2}\right)\rho^2 d\rho} \\
\frac{\partial u(\vec{r}\,)}{\partial y_0}=\frac{2\pi \textrm{NA}}{\lambda}\frac{y}{\|\vec{r}\,\|} \int_{0}^{1}{J_1\left(\frac{2\pi \textrm{NA}}{\lambda} \rho \|\vec{r}\,\|\right)\cos\left({\frac{\pi \textrm{NA}^2 z_0}{\lambda n_m}\rho^2}\right)\rho^2 d\rho} \\
\frac{\partial v(\vec{r}\,)}{\partial y_0}=\frac{2\pi \textrm{NA}}{\lambda}\frac{y}{\|\vec{r}\,\|} \int_{0}^{1}{J_1\left(\frac{2\pi \textrm{NA}}{\lambda} \rho \|\vec{r}\,\|\right)\sin\left({\frac{\pi \textrm{NA}^2 z_0}{\lambda n_m}\rho^2}\right)\rho^2 d\rho} \\
\end{gather*}
\begin{gather*}
\frac{\partial u(\vec{r}\,)}{\partial \lambda}=\frac{2\pi \textrm{NA}}{\lambda^2}\|\vec{r}\,\| \int_{0}^{1}{J_1\left(\frac{2\pi \textrm{NA}}{\lambda} \rho \|\vec{r}\,\|\right)\cos\left({\frac{\pi \textrm{NA}^2 z_0}{\lambda n_m}\rho^2}\right)\rho^2 d\rho} \\
+\frac{\pi \textrm{NA}^2 z_0}{\lambda^2 n_m} \int_{0}^{1}{J_0\left(\frac{2\pi \textrm{NA}}{\lambda} \rho \|\vec{r}\,\|\right)\sin\left({\frac{\pi \textrm{NA}^2 z_0}{\lambda n_m}\rho^2}\right)\rho^3 d\rho}\\
\frac{\partial v(\vec{r}\,)}{\partial \lambda}=\frac{2\pi \textrm{NA}}{\lambda^2}\|\vec{r}\,\| \int_{0}^{1}{J_1\left(\frac{2\pi \textrm{NA}}{\lambda} \rho \|\vec{r}\,\|\right)\sin\left({\frac{\pi \textrm{NA}^2 z_0}{\lambda n_m}\rho^2}\right)\rho^2 d\rho} \\
-\frac{\pi \textrm{NA}^2 z_0}{\lambda^2 n_m} \int_{0}^{1}{J_0\left(\frac{2\pi \textrm{NA}}{\lambda} \rho \|\vec{r}\,\|\right)\cos\left({\frac{\pi \textrm{NA}^2 z_0}{\lambda n_m}\rho^2}\right)\rho^3 d\rho}
\end{gather*}
\begin{gather*}
\frac{\partial A}{\partial \lambda}= \int_{\mathbb{R}^2} { \left(2u(\vec{r}\,)\frac{\partial u(\vec{r}\,)}{\partial \lambda} + 2v(\vec{r}\,)\frac{\partial v(\vec{r}\,)}{\partial \lambda}\right) d\vec{r}} \\
\end{gather*}
\begin{align*}
\frac{\partial w(\vec{r}-\vec{r}_0)}{\partial x_0}&=\int_{0}^{1}{A \frac{\partial}{\partial x_0}\left[J_0\left(\frac{2\pi \textrm{NA}}{\lambda} \rho \|\vec{r}-\vec{r}_0\|\right)\right]e^{j\frac{\pi \textrm{NA}^2 z_0}{\lambda n_m}\rho^2}\rho d\rho} \\
&=-\frac{2\pi \textrm{NA}}{\lambda}\frac{x-x_0}{\|\vec{r}-\vec{r}_0\|}\int_{0}^{1}{A J_0'\left(\frac{2\pi \textrm{NA}}{\lambda} \rho \|\vec{r}-\vec{r}_0\|\right)e^{j\frac{\pi \textrm{NA}^2 z_0}{\lambda n_m}\rho^2}\rho^2 d\rho}\\
\frac{\partial w(\vec{r}-\vec{r}_0)}{\partial y_0}&=-\frac{2\pi \textrm{NA}}{\lambda}\frac{y-y_0}{\|\vec{r}-\vec{r}_0\|}\int_{0}^{1}{A J_0'\left(\frac{2\pi \textrm{NA}}{\lambda} \rho \|\vec{r}-\vec{r}_0\|\right)e^{j\frac{\pi \textrm{NA}^2 z_0}{\lambda n_m}\rho^2}\rho^2 d\rho}\\
\frac{\partial w(\vec{r}-\vec{r}_0)}{\partial \lambda}&=\int_{0}^{1}{A \left( \frac{\partial}{\partial \lambda}\left[J_0\left(\frac{2\pi \textrm{NA}}{\lambda} \rho \|\vec{r}-\vec{r}_0\|\right)\right]e^{j\frac{\pi \textrm{NA}^2 z_0}{\lambda n_m}\rho^2} +
J_0\left(\frac{2\pi \textrm{NA}}{\lambda} \rho \|\vec{r}-\vec{r}_0\|\right)\frac{\partial}{\partial \lambda}\left[e^{j\frac{\pi \textrm{NA}^2 z_0}{\lambda n_m}\rho^2}\right]\right)\rho d\rho} \\
&=-\frac{2\pi \textrm{NA}}{\lambda^2}\|\vec{r}-\vec{r}_0\|\int_{0}^{1}
{A J_0'\left(\frac{2\pi \textrm{NA}}{\lambda} \rho \|\vec{r}-\vec{r}_0\|\right)e^{j\frac{\pi \textrm{NA}^2 z_0}{\lambda n_m}\rho^2} \rho^2 d\rho}\\
&-\frac{\pi \textrm{NA}^2 z_0}{\lambda^2 n_m} \int_{0}^{1}{ A J_0 \left(\frac{2\pi \textrm{NA}}{\lambda} \rho \|\vec{r}-\vec{r}_0\|\right)
e^{j\frac{\pi \textrm{NA}^2 z_0}{\lambda n_m}\rho^2}
\rho^3 d\rho}
\end{align*}

## Slide 17: CRLB: Circular Aperture Case
For z=0 (perfect focusing)
Notes: \begin{gather*}
q_{z_0}(\vec{r}-\vec{r}_0\,)=\left\|A \int_{0}^{1}{J_0\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}-\vec{r}_0\| \rho\right)e^{-j\frac{\pi}{\lambda}\rho^2 z_0\left(\frac{\textrm{NA}^2}{n_i}\right) }\rho d\rho} \right\|^2\\
A\textsf{: normalization constant,}\: n_i\textsf{: refractive index of the immersion oil}\\
J_0\textsf{: zeroth order Bessel function of the first kind}
\end{gather*}
\begin{gather*}
q_{z_0}(\vec{r}-\vec{r}_0\,)=\frac{1}{A_{z_0}}\left(u_{z_0}^2(\vec{r}-\vec{r}_0\,) + v_{z_0}^2(\vec{r}-\vec{r}_0\,)\right)\\
u_{z_0}(\vec{r}-\vec{r}_0\,)=\mathlarger{\int}_{0}^{1}{J_0\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}-\vec{r}_0\| \rho\right)\cos\left({\frac{\pi}{\lambda}\rho^2 z_0 \left(\frac{\textrm{NA}^2}{n_i}\right)}\right)\rho d\rho} \\
v_{z_0}(\vec{r}-\vec{r}_0\,)=\mathlarger{\int}_{0}^{1}{J_0\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}-\vec{r}_0\| \rho\right)\sin\left({\frac{\pi}{\lambda}\rho^2 z_0 \left(\frac{\textrm{NA}^2}{n_i}\right)}\right)\rho d\rho} \\
A_{z_0} = \int_{\mathbb{R}^2} {\left(u^2(\vec{r}\,) + v^2(\vec{r}\,)\right)}d\vec{r}
\end{gather*}

## Slide 18: Impact of z variation
Defocusing causes biasing for wavelength estimation
  Accuracy bound should take into account both bias & CRLB
Generalized CRLB in presence of bias
CRLB is z dependent
Notes: \begin{gather*}
q_{z_0}(\vec{r}-\vec{r}_0\,)=\left\|A \int_{0}^{1}{J_0\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}-\vec{r}_0\| \rho\right)e^{-j\frac{\pi}{\lambda}\rho^2 z_0\left(\frac{\textrm{NA}^2}{n_i}\right) }\rho d\rho} \right\|^2\\
A\textsf{: normalization constant,}\: n_i\textsf{: refractive index of the immersion oil}\\
J_0\textsf{: zeroth order Bessel function of the first kind}
\end{gather*}
\begin{gather*}
q_{z_0}(\vec{r}-\vec{r}_0\,)=\frac{1}{A_{z_0}}\left(u_{z_0}^2(\vec{r}-\vec{r}_0\,) + v_{z_0}^2(\vec{r}-\vec{r}_0\,)\right)\\
u_{z_0}(\vec{r}-\vec{r}_0\,)=\mathlarger{\int}_{0}^{1}{J_0\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}-\vec{r}_0\| \rho\right)\cos\left({\frac{\pi}{\lambda}\rho^2 z_0 \left(\frac{\textrm{NA}^2}{n_i}\right)}\right)\rho d\rho} \\
v_{z_0}(\vec{r}-\vec{r}_0\,)=\mathlarger{\int}_{0}^{1}{J_0\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}-\vec{r}_0\| \rho\right)\sin\left({\frac{\pi}{\lambda}\rho^2 z_0 \left(\frac{\textrm{NA}^2}{n_i}\right)}\right)\rho d\rho} \\
A_{z_0} = \int_{\mathbb{R}^2} {\left(u^2(\vec{r}\,) + v^2(\vec{r}\,)\right)}d\vec{r}
\end{gather*}

## Slide 19: Impact of z variation
Magnification factor & CRLB vs. z
  N=1500, λ=600nm, NA=1.4, ni=1.5, Born & Wolf model
Notes: \begin{gather*}
q_{z_0}(\vec{r}-\vec{r}_0\,)=\left\|A \int_{0}^{1}{J_0\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}-\vec{r}_0\| \rho\right)e^{-j\frac{\pi}{\lambda}\rho^2 z_0\left(\frac{\textrm{NA}^2}{n_i}\right) }\rho d\rho} \right\|^2\\
A\textsf{: normalization constant,}\: n_i\textsf{: refractive index of the immersion oil}\\
J_0\textsf{: zeroth order Bessel function of the first kind}
\end{gather*}
\begin{gather*}
q_{z_0}(\vec{r}-\vec{r}_0\,)=\frac{1}{A_{z_0}}\left(u_{z_0}^2(\vec{r}-\vec{r}_0\,) + v_{z_0}^2(\vec{r}-\vec{r}_0\,)\right)\\
u_{z_0}(\vec{r}-\vec{r}_0\,)=\mathlarger{\int}_{0}^{1}{J_0\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}-\vec{r}_0\| \rho\right)\cos\left({\frac{\pi}{\lambda}\rho^2 z_0 \left(\frac{\textrm{NA}^2}{n_i}\right)}\right)\rho d\rho} \\
v_{z_0}(\vec{r}-\vec{r}_0\,)=\mathlarger{\int}_{0}^{1}{J_0\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}-\vec{r}_0\| \rho\right)\sin\left({\frac{\pi}{\lambda}\rho^2 z_0 \left(\frac{\textrm{NA}^2}{n_i}\right)}\right)\rho d\rho} \\
A_{z_0} = \int_{\mathbb{R}^2} {\left(u^2(\vec{r}\,) + v^2(\vec{r}\,)\right)}d\vec{r}
\end{gather*}

## Slide 20: Impact of z variation
Assuming certain distribution for z, performance bound with z variation can be calculated
  Error for each z: CRLB(z) + bias2(z)
  Final CRLB is calculated as a weighted sum
  For M(z)=1, converges to the CRLB @ z=0
  For Δzmax=100nm & uniform distribution, CRLB=9.86nm (originally 4.97nm)
  For Δzmax=500nm, CRLB=177nm!
Notes: \begin{gather*}
q_{z_0}(\vec{r}-\vec{r}_0\,)=\left\|A \int_{0}^{1}{J_0\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}-\vec{r}_0\| \rho\right)e^{-j\frac{\pi}{\lambda}\rho^2 z_0\left(\frac{\textrm{NA}^2}{n_i}\right) }\rho d\rho} \right\|^2\\
A\textsf{: normalization constant,}\: n_i\textsf{: refractive index of the immersion oil}\\
J_0\textsf{: zeroth order Bessel function of the first kind}
\end{gather*}
\begin{gather*}
q_{z_0}(\vec{r}-\vec{r}_0\,)=\frac{1}{A_{z_0}}\left(u_{z_0}^2(\vec{r}-\vec{r}_0\,) + v_{z_0}^2(\vec{r}-\vec{r}_0\,)\right)\\
u_{z_0}(\vec{r}-\vec{r}_0\,)=\mathlarger{\int}_{0}^{1}{J_0\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}-\vec{r}_0\| \rho\right)\cos\left({\frac{\pi}{\lambda}\rho^2 z_0 \left(\frac{\textrm{NA}^2}{n_i}\right)}\right)\rho d\rho} \\
v_{z_0}(\vec{r}-\vec{r}_0\,)=\mathlarger{\int}_{0}^{1}{J_0\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}-\vec{r}_0\| \rho\right)\sin\left({\frac{\pi}{\lambda}\rho^2 z_0 \left(\frac{\textrm{NA}^2}{n_i}\right)}\right)\rho d\rho} \\
A_{z_0} = \int_{\mathbb{R}^2} {\left(u^2(\vec{r}\,) + v^2(\vec{r}\,)\right)}d\vec{r}
\end{gather*}

## Slide 21: 3D + Color: Multifocal Plane
Total Fisher information matrix is the sum of FI from each imaging plane
Biplane setting
  Ntotal=1500, λ=600nm, NA=1.4, ni=1.5
  50/50 beam split (750 photons/plane)
  Distance between planes: 500nm
Plane 1
Plane 2
Notes: \begin{gather*}
q_{z_0}(\vec{r}-\vec{r}_0\,)=\left\|A \int_{0}^{1}{J_0\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}-\vec{r}_0\| \rho\right)e^{-j\frac{\pi}{\lambda}\rho^2 z_0\left(\frac{\textrm{NA}^2}{n_i}\right) }\rho d\rho} \right\|^2\\
A\textsf{: normalization constant,}\: n_i\textsf{: refractive index of the immersion oil}\\
J_0\textsf{: zeroth order Bessel function of the first kind}
\end{gather*}
\begin{gather*}
q_{z_0}(\vec{r}-\vec{r}_0\,)=\frac{1}{A_{z_0}}\left(u_{z_0}^2(\vec{r}-\vec{r}_0\,) + v_{z_0}^2(\vec{r}-\vec{r}_0\,)\right)\\
u_{z_0}(\vec{r}-\vec{r}_0\,)=\mathlarger{\int}_{0}^{1}{J_0\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}-\vec{r}_0\| \rho\right)\cos\left({\frac{\pi}{\lambda}\rho^2 z_0 \left(\frac{\textrm{NA}^2}{n_i}\right)}\right)\rho d\rho} \\
v_{z_0}(\vec{r}-\vec{r}_0\,)=\mathlarger{\int}_{0}^{1}{J_0\left(\frac{2\pi}{\lambda}\textrm{NA} \|\vec{r}-\vec{r}_0\| \rho\right)\sin\left({\frac{\pi}{\lambda}\rho^2 z_0 \left(\frac{\textrm{NA}^2}{n_i}\right)}\right)\rho d\rho} \\
A_{z_0} = \int_{\mathbb{R}^2} {\left(u^2(\vec{r}\,) + v^2(\vec{r}\,)\right)}d\vec{r}
\end{gather*}

## Slide 22: 3D + Color: Double Helix/Astigmatic
Smith, C. et al., Simultaneous measurement of emission color and 3D position of single molecules. Optics Express (2016).
Notes: S_i(f) = \frac{{i_0}^2}{2}\left[ \frac{{\tau}_c/(2\pi)}{1+\left(\pi f {\tau}_c \right)^2} \left\{ 1- e^{-\frac{2\tau}{{\tau}_c}} \left[ \cos(2\pi f\tau) + \frac{1}{\pi f \tau_c} \sin(2\pi f \tau) \right] \right\} \right]
\tau_c = \frac{2}{\textrm{linewidth}},\quad i_0: \textrm{nominal photocurrent amplitude},\quad \tau: \textrm{path delay}

## Slide 23: 3D + Color: Double Helix/Astigmatic
Smith, C. et al., Simultaneous measurement of emission color and 3D position of single molecules. Optics Express (2016).
Notes: S_i(f) = \frac{{i_0}^2}{2}\left[ \frac{{\tau}_c/(2\pi)}{1+\left(\pi f {\tau}_c \right)^2} \left\{ 1- e^{-\frac{2\tau}{{\tau}_c}} \left[ \cos(2\pi f\tau) + \frac{1}{\pi f \tau_c} \sin(2\pi f \tau) \right] \right\} \right]
\tau_c = \frac{2}{\textrm{linewidth}},\quad i_0: \textrm{nominal photocurrent amplitude},\quad \tau: \textrm{path delay}

## Slide 24: To be added
CRLB including EMCCD readout noise
Fluorophore spectrum dependence
Actual PSF1 estimation of existing microscope
Compared to the fitting simulation dataset (with MLE & LS)
…
[1] Theer, P. et al., PSFj: know your fluorescence microscope. Nat. Meth. (2014).
Notes: S_i(f) = \frac{{i_0}^2}{2}\left[ \frac{{\tau}_c/(2\pi)}{1+\left(\pi f {\tau}_c \right)^2} \left\{ 1- e^{-\frac{2\tau}{{\tau}_c}} \left[ \cos(2\pi f\tau) + \frac{1}{\pi f \tau_c} \sin(2\pi f \tau) \right] \right\} \right]
\tau_c = \frac{2}{\textrm{linewidth}},\quad i_0: \textrm{nominal photocurrent amplitude},\quad \tau: \textrm{path delay}

## Slide 25: Theoretical Performance Analysis: Possible Source of Bias

## Slide 26: Possible Bias Sources
Chromatic/spherical aberration
Spectral profile of the fluorophore (bandwidth, skewness)
Non-uniform background
Deconvolution artifacts
Motion blur
Notes: S_i(f) = \frac{{i_0}^2}{2}\left[ \frac{{\tau}_c/(2\pi)}{1+\left(\pi f {\tau}_c \right)^2} \left\{ 1- e^{-\frac{2\tau}{{\tau}_c}} \left[ \cos(2\pi f\tau) + \frac{1}{\pi f \tau_c} \sin(2\pi f \tau) \right] \right\} \right]
\tau_c = \frac{2}{\textrm{linewidth}},\quad i_0: \textrm{nominal photocurrent amplitude},\quad \tau: \textrm{path delay}

## Slide 27: Experiment Design

## Slide 28: Expected Impact of Experiment Settings
NA
Index matching
Magnification
Fluorophore density
Fluorophore spectrum
Illumination
Software setting
Notes: S_i(f) = \frac{{i_0}^2}{2}\left[ \frac{{\tau}_c/(2\pi)}{1+\left(\pi f {\tau}_c \right)^2} \left\{ 1- e^{-\frac{2\tau}{{\tau}_c}} \left[ \cos(2\pi f\tau) + \frac{1}{\pi f \tau_c} \sin(2\pi f \tau) \right] \right\} \right]
\tau_c = \frac{2}{\textrm{linewidth}},\quad i_0: \textrm{nominal photocurrent amplitude},\quad \tau: \textrm{path delay}

## Slide 29: Backup Slides

## Slide 31: Theoretical Performance Evaluation
Vectorial model1 (considers dipole orientation/polarization)
[1] Mortensen, K. I. et al., Optimized localization analysis for single-molecule tracking and super-resolution microscopy. Nat. Meth. (2010).
Notes: \begin{gather*}
w(\vec{r}\,)=\frac{1}{w_n}\int_{|\vec{\rho}\,|\leq 1}{d^2\rho A(\vec{\rho}\,)\exp{\left(\frac{2\pi i W(\vec{\rho}\,)}{\lambda}\right)}}q(\vec{\rho}\,)\exp{\left(-i \vec{k}(\vec{\rho}\,)\cdot \vec{r}\right)}\\
\vec{k}(\vec{\rho}\,)=\frac{2\pi}{\lambda}\left(\textrm{NA}\rho_x,\textrm{NA}\rho_y,\sqrt{n^2-\textrm{NA}^2\vec{\rho}\,^2}\right)\\
\vec{\rho}\,\textsf{: pupil plane coordinate,}\:W(\vec{\rho}\,)\textsf{: pupil plane modulation mask,}\:\vec{k}(\vec{\rho}\,)\textsf{: wavevector}
\end{gather*}
\begin{align*}
\frac{\partial\mu_k}{\partial\theta}&=\frac{2N}{3}\int_{\mathscr{D}_k}{dxdy\,\textrm{Re}\left\{w(\vec{r}-\vec{r}_0)\frac{\partial w(\vec{r}-\vec{r}_0)}{\partial \theta}\right\}}\\
\frac{\partial w(\vec{r}-\vec{r}_0)}{\partial \vec{r}_0}&=\frac{\partial w(\vec{r}-\vec{r}_0)}{\partial x_0}=\frac{\partial w(\vec{r}-\vec{r}_0)}{\partial y_0}=\frac{\partial w(\vec{r}-\vec{r}_0)}{\partial z_0}\\
&=\frac{i}{w_n}\int_{|\vec{\rho}\,|\leq 1}{d^2\rho A(\vec{\rho}\,)\exp{\left(\frac{2\pi i W(\vec{\rho}\,)}{\lambda}\right)}}q(\vec{\rho}\,)\vec{k}(\vec{\rho}\,)\exp{\left(-i \vec{k}(\vec{\rho}\,)\cdot (\vec{r}-\vec{r}_0)\right)}\\
\frac{\partial w(\vec{r}-\vec{r}_0)}{\partial \lambda}&=-\frac{2\pi i}{w_n \lambda ^2}\int_{|\vec{\rho}\,|\leq 1}{d^2\rho A(\vec{\rho}\,)\exp{\left(\frac{2\pi i W(\vec{\rho}\,)}{\lambda}\right)}}q(\vec{\rho}\,)W(\vec{\rho}\,)\exp{\left(-i \vec{k}(\vec{\rho}\,)\cdot (\vec{r}-\vec{r}_0)\right)}\\
&+\frac{1}{\lambda}(\vec{r}-\vec{r}_0)\cdot\frac{\partial w(\vec{r}-\vec{r}_0)}{\partial x_0}
\end{align*}